Microfluidics 110Y and 110L

Materials

- E. coli sample in ice

- 200ml 70 % EtOH

- 3 liter H2O

- Microfluidizer

- Glass tube

Procedure

- Connect high pressure system to microfluidizer. Connect glass tube to sample loading position. Connect plastic tube to waste.

- Check microfluidizer is turn off. Add H2O to glass tube. Turn on high pressure system pump slowly.

- Wash microfluidizer chamber three times with H2O. Wash microfluidizer chamber once with buffer. Cool microfluidizer chamber with ice.

- Load sample and break cell 3-5 times.

- Wash microfluidizer chamber three times with H2O. Wash microfluidizer chamber once with 70 % EtOH.

- Keep 50ml 70 % EtOH in glass tube. Turn off high pressure system. Turn on microfluidizer to release pressure and then turn off it.

- Dissolve ice with H2O.

- Separate high pressure system with microfluidizer. Collect plastic tube well. Cover sample loading position with aluminum foil.

- Clean up microfluidizer and keep dry.